Antenna beam pattern evaluation of ALMA band 9 subsystem 
at North American Integration Center

Work visit technical report

A. Baryshev

13 June 2008

Introduction

This report describes band 9 beam pattern measurements which were made at NA ALMA front end integration center during the visit of ALMA band 9 group representatives from 2nd to 7th of June 2008. The Band 9 cartridge system was mounted within ALMA front end cryostat #2 and the overall system included the cryostat window and filters. Measurement set-up, experiment condition and sequence and results of beam analysis are described. After iterative improvements, the NA FEIC measurement system was found to be adequate for beam characterization in band 9 frequency range. The measured band 9 optics beam was found to be in good agreement with measurements done by band 9 cartridge manufacturer and total aperture efficiency was found to satisfy ALMA requirements.

Set-up and measurement system description

Band 9 cartridge beam forming optics is all contained at the 4K temperature level within the Front-End cryostat. The ALMA specified entrance window (see fig. 1) and infrared filters were used in the system and the cartridge was cooled to the nominal temperature levels to operate the receiver. During the whole session receiver 4 K stage was approx at 3.6 K and no temperature stabilizing by heaters was applied, likewise for cartridge measurements. The Band 3, 6 and 7 cartridge systems were mounted in the same front-end but were kept powered down. ALMA IF switch and WCA was part of room temperature IF part which followed by homodyne phase and amplitude detection IF breadboard design of which is described in [1]. This design is conceptually similar to the one used by band 9 cartridge manufacturers to characterize the beam. An LO warm multiplier part of WCA ending by a quintupler multiplier stage was used as signal source. It was mounted on a 1 x 1 x 0.1 m mechanical scanner with possibility to rotate it around beam propagation axis by 90 degrees for cross polarization measurements. Reference channel makes use of leakage signal from photonic reference port of the cartridge WCA. Both source and main cartridge WCA’s were using signal synthesized sweepers to provide reference frequencies. A flared horn has been mounted to a quintupler stage to form wide test beam which is adequate for measuring f#8 cartridge output beam.

A laser interferometer stage [2] was used to align scanning plane with respect to antenna mounting interface of the front-end. Absolute measurement accuracy of approx 7 microns can be achieved on all axes.

Initial status

Before the measurements commenced, NA FEIC has done preliminary measurement of band 9 cartridge beam. The raw test results and far field beam, derived directly from raw results without correction for probe beam pattern is shown in Fig 2-4 for 661 GHz signal frequency. After clearing a communication issues about data format, it was possible to arrive to an aperture efficiency numbers of order of 80% in the optimal beam direction and optimal focus positions.

Fig. 1, Photograph of ALMA Front-End cryostat with band 9 entrance window indicated by an arrow

Fig. 2, Amplitude plot of initial NA FEIC data	Fig. 3, Phase plot of initial NA FEIC data

Fig. 4, Far field plot of initial NA FEIC data	Fig. 5, Far field plot of initial NA FEIC data

The beam pointing angles were calculated by a center of mass formula and are (0.16,-1.34) degrees which should compare to nominal angle (0, -0.94) degrees for band 9 cartridge.

Aperture efficiency numbers evaluated using the same data by NA FEIC group and band 9 group coincide within 0.1 percent when same input parameters are used.

Mechanical stage stability and alignment has been thoroughly investigated by using laser interferometer measuring system. Necessary correction coefficients have been entered into XYZ stages control software.

Initial analysis and improvement steps

Scanner movements

It is apparent from fig. 2, that there is a problem with backlash in the current system, which can be mechanical (scanner position backlash) or through cable, attached to moving stage. It was found that scanner was moving in meandering (fig 6 a) way which can explain vertical modulation of contour levels in fig 2, 3 and artificial side-lobes at +- 15 deg Y angle in fig. 4

a)						b)

Fig. 6, Meandering a) and one side approach b) scanning sequences. Dashed lines indicate idle scanner movements

NSI software can be configured to do a one side approach scanning fig. 6 b) which takes about only 10% longer because of high movement speeds when not taking data. Once this approach is used the above mentioned artifacts due to backlash disappear (see for instance fig. 7). All following measurements were done using approach of 6 b)

Probe geometry

Fig. 5 show central part of the beam in far field. Contours of secondary antenna reflector are shown by white dashed line around beam center found by a center of mass algorithm. Already adequate and close to expected edge illumination level of -12 dB is detected. This means that probe used by NA FEIC has adequately wide beam and does not lead to significant loss in measurement quality. In order to check this fact experimentally an alternative flared horn probe has been brought to NA FEIC (see fig 7 Andrey add). Apparent illumination edge taper has increased, indicating that it is important to use correct probe, or compensate the measured data for probe beam pattern before evaluating aperture efficiency. It can be shown that the flared horn probe, which was used, does not require significant correction of the data.

Phase stability

Phase stability is important parameter of measurement set-up which influences apparent beam pointing direction detected from the measured data. It was found during test run, that pointing repeatability of setup was about 0.5 degrees. After careful investigation, a fault in one of the flexible phase cables was determined, the cable was replaced and rerouted in order to give the least phase error during scanning. After this a pointing repeatability of 0.02 degrees have been achieved and verified by consecutive scans.

Standing wave compensation

As it was demonstrated by other groups, it is possible to compensate standing waves between probes and receiver by making scan of two planes a1 and a2, separated by lambda/4 distance. By careful comparison of NSI software data output, it was found that it erroneously uses average (a1+a2)/2 instead of (a1+I*a2)/2 or (a1-I*a2)/2 depending on the sideband used in detection. It was demonstrated, that despite of the fact that a2 data is taken 40 min after a1, phase and amplitude stability of setup allows for efficient standing wave compensation. It was also apparent, that by using correct probe and absorbers geometry (which were adequate from the beginning) it is not always necessary to use two plane measurements to characterize the cartridge beam adequately.

Measurement sequence and analysis

After performance of measurement set-up was verified, a set of beam measurements has been performed to make the final assessment of the beam quality of band 9.

First probe was rotated by 45 degrees and two consecutive measurements were done for polarization 0 and 1 switching over between polarizations by reconnecting IF signal. This set is done to verify beam pointing, aperture efficiency and beam squint.

Insert analysis

Cross and co-polar beam pattern have been measured for one polarization by rotating probe 90 degrees without changing output level.

Insert analysis

Since measurements of beam pattern were interleaved in time with cartridge noise measurements, it was no time available to do more detailed measurements. The existing measurements provide solid evidence that measurements of cartridge at NA FEIC and at cartridge manufacturer are consistent, and band 9 front-end meets most of optics specifications.

Conclusions

During this debug and measurement session, it was demonstrated that:

- NA FEIC measurement beam pattern set-up is adequate for ALMA front-end optics specifications verification, provided proper measurement method is chosen

- NA FEIC and band 9 group analysis methods were cross checked arriving to very similar aperture efficiency data using the same data set

- Band 9 cartridge measurements done within front-end and by manufacturer match each other and band 9 cartridge meets optics specification (with exclusion of cross polarization) as demonstrated by cartridge manufacturer measurements

Next steps

- Firm, spelled out cartridge optics verification procedure needs to be developed at (NA) FEIC that includes: beam measurement sequence (plan), and data analysis procedures, agreed with all involved parties

- Further improvement of measurement set-up is possible by getting better phase stable cable. Band 9 group to provide specifications of their hardware.

- Beam measurements for several frequencies and polarization channels can be automated using scripts and appropriate hardware. This will allow to save time and use night time for beam scanning efficiently.  It is already ongoing.

- It is recommended to establish working group which will facilitate experience exchange within the ALMA project and improve its problem solving capabilities

Acknowledgements

I would like to thank G.A.Ediss, M.McLeod and R.Williams for their assistance during my visit to the NA FEIC.

References

[1] NA FEIC test and measurement system design. 2005-07-27. FEND-40.09.03.00-002-A-DSN and

Modifications to NA FEIC beam pattern measuring system. 2007-03-21. FEND-40.09.03.00-033-A-REP

[2] FARO laser tracker Xi, www.faro.com